## Slide 1
W.R. Goodman1, M.D. Nornberg1, A. Wolfmeister1, A.M. Wright1, B. Geiger1, and the HSX Team
1HSX Plasma Laboratory, University of Wisconsin, Madison, Wisconsin, USA
Commissioning a High-Repetition Rate, 1 Joule Nd:YAG Laser for the HSX Thomson Scattering System
We report on the installation and commissioning of a major TS upgrade at the HSX stellarator, centered on a new Nd:YAG laser capable of burst operation of 80 consecutive 1 J, 10 ns pulses at 400 Hz. A 2 kHz mode supports up to 10 pulses per burst, limited by thermal loading. The upgrade includes integrating the laser with HSX beam-delivery optics and shot-to-shot pulse-energy monitoring to verify stability. To accommodate the increased data rate, a high-speed digitizer (3.2 GSa/s) and low-noise preamplifiers, providing an ~8.6 dB improvement in signal-to-noise ratio over the legacy digitizer chain [1,2]. In the near term, increasing the time resolution of the TS system will enable the study of plasma breakdown effects, MHD effects, impurity injection behavior, and time-resolved density profile adjustments.
This work is supported by US DOE award DE-FG02-93ER54222
Acknowledgements
26th Topical Conference on High Temperature Plasma Diagnostics. June 7 – June 11, 2026. Cambridge, MA
[1] K. Zhai et al, Rev. Sci. Instrum., 75, 3900-3902 (2004)
[2] W.R. Goodman et al, Rev. Sci. Instrum., 94, 093507 (2024)
[3] W. R. Goodman, B. Geiger, Rev. Sci. Instrum., 95, 093507 (2024)
[4] W. R. Goodman, GitHub Repo, doi.org/10.5281/zenodo.20269598 (2026)
The HSX Stellarator
Major radius: 1.20 m Average minor radius: 0.12 m Plasma volume: 0.44 m3 Field periods: 4 Number of main coils: 48 Number of auxiliary coils: 48
HSX TS System
10 spatial channels [1] 1.6*20 mm scattering area 0.01 ≤ Te ≤ 4 keV 0.5 ≤ ne ≤ 10 [1018 m-3 ] 78.4 MHz detection BW 3.2 GSa/s digitization
TSSM Filters[3]
Machine learning optimized dual-passband filter set. Reduces error and/or increases bandwidth by more efficiently using the collected photons from a Thomson scattering event.
Beam Path
Even with high-reflectivity mirrors, each interaction results in a fractional power loss. Laser was mounted in a manner that there is only one optic interaction along the beam path.
Beam Dump
A Coherent J-50MB-YAG-1528 Energy Monitor is used as a combination beam dump and shot-to-shot energy measurement system enabling individual energy measurements for each TS laser pulse.
Motivation
Thomson Scattering on HSX
References
An InnoLas SpitLight DPSS EVO IV has been incorporated into daily operations on the HSX Thomson scattering system. Mean beam energy through vessel = , or of initial laser output.
Results
InnoLas Laser System
SpitLight DPSS EVO IV Nd:YAG ~1 Joule, 1064 nm, 8 ns pulse 100 Hz continuous operation 80 pulses @ 400 Hz, 4 x 10 pulses @ 2 kHz Cold-rod operation standard
Laser Mounting
The EVO IV was mounted on an optical table ~1 m above the main TS optical table, eliminating at least two optical elements required for a beam periscope. Shown are a -waveplate (QWP), a -waveplate (HWP), and an interlocking beam shutter on the EVO IV exit.
Beam Energy
Beam energies were measured with a Coherent J-50MB-YAG-1528 energy sensor (100 Hz / 400 Hz) or a Coherent J-50MT-10KHZ energy sensor (2 kHz). The SpitLight IV is capable of continuous operation of 100 Hz. 100 Hz operation has a short warm-up period of ~300 pulses before leveling off at maximum energy output. Burst mode operation is enabled for up to 80 pulses at 400 Hz or four sets of ten pulses at 2 kHz. Burst mode operation does not have the initial warm-up period and does not have the same thermal lensing losses seen in 100 Hz operation.
Beam Polarization
Referring to the image of the HSX TS beam dump, a ~6” Brewster window is used as the beam exit. The Brewster window is installed so that p-polarized light is aligned with the magnetic axis, , which aids in polarization alignment. The one mirror that the HSX TS beam interacts with is at a non-45-degree angle, introducing ellipticity into the beam polarization. Polarization is corrected by adjusting a QWP and HWP in the beam path with the goal of maximizing measured energy through the Brewster window.
Final settings: waveplate @ 82° and waveplate @ 18°.
Uncorrected beam and phase shift
Energy Through Vessel
Warm-rod vs. Cold-rod Operation
When a diode-pumped solid-state laser operates with cold rods the thermal lens effect within the gain medium is minimal, meaning the rod behaves closer to a flat optical element, which can result in a larger or poorly matched mode volume. As the rods warm under sustained pumping, a roughly parabolic radial temperature gradient develops due to waste heat, creating a positive thermal lens that effectively shortens the cavity's equivalent focal length. A well-designed resonator is typically optimized for this steady-state thermal lens, so warm-rod operation yields a stable, well-defined mode, predictable beam divergence, and consistent output power. For example, for 100 Hz continuous operation, the InnoLas laser system incorporates a warm-up period before the laser functions at the factory design point.
Cold-rod measurements
Warm-rod measurements Arduino GIGA R1 needed to run warm-rod code.
Laser Implementation & Results
For access to all data, materials, and MATLAB functions used to generate this presentation, including code to enable warm-rod operation on an InnoLas laser, please scan the QR code or go to https://github.com/goodystacks/htpd2026[4].
For more info on HSX, please visit https://hsx.wisc.edu/. For more info on the University of Wisconsin or the Nuclear Engineering & Engineering Physics department, visit engineering.wisc.edu/ departments/nuclear-engineering- engineering-physics/
Info & Resources
Warm-up Period
Commission 2 kHz burst mode for daily operation
Design an ultra-low reflection beam expander lens system to eliminate possible stray light and avoid damage to the Coherent EnergyMax-USB J-50MT-10KHZ Energy Sensor
Commission warm-rod mode for daily operation
Add control channels to the HSX control system, verify operational stability, reperform Raman alignment and calibration.
Future Work